## Abstract

Under suitable conditions it is possible to employ very simple probes with the MoisturePoint MP-917 readout unit. These simple probes may be user-fabricated at very low cost. This technical brief describes how to use some of the diagnostic capabilities of the MP-917 to evaluate the operation of the proposed probe in the soil to be examined, how to determine the probe calibration factors, and shows an example of probe fabrication.

Moisture•PointTM

MP-917 Technical Brief

## Table of Contents

µIntroduction	3

Connection of Probe to MP-917	4

Diode Waveform	4

Open Reflection Waveform	5

Setting 'Instrument Default' Parameters	6

Setting 'Probe Default' Parameters	9

Typical Automatic-Operation Waveforms	10

Appendix 1 - Probe Construction	11

Appendix 2 - Probe Calibration	13

Introductiontc "Introduction"\l§

The MP-917 measures the round-trip propagation time of a pulse along a probe segment in two steps.  First, it measures the round-trip time to the beginning of the segment--which is marked by the presence of a diode.  Second, it measures the round-trip time to the end of the segment.  The difference between these two times is the round-trip propagation time along the segment.  Usually, the end of the segment being measured is also marked with a diode, but in the case of the single-diode probes, the second end of the segment is just the open end of the probe.

The use of diodes to define the segment boundaries allows the controller to create a distinctive reflection of the transmitted pulse when the pulse reaches the diode.  In fact, the diode is turned on and off during the measurement and the instrument observes the difference between the signal on the probe transmission line with the diode on and off.  Ideally, the only difference between the on and off waveform occurs at the time corresponding to the location of the diode. Selecting a diode first at the beginning of a segment and then at the end of the segment allows the times to these two segment ends to be obtained with the greatest precision. (See Hook et al. , "Remote Diode Shorting Improves Measurement of Soil Water by Time Domain Reflectometry" Soil Sci. Soc. Am. J. 56:1384-1391 (1992).)

However, when a short probe is placed in fairly well-behaved soil (not too conductive) the transmit pulse is reflected from the open end strongly enough, and reflected with a distinctive waveshape, so that the MP-917 may use this open-end reflection to locate the end of the probe. Therefore, a single-segment probe may be fabricated which has only a diode located at the top of the probe, resulting in a simpler overall construction which a user may undertake himself.

The use of these single-diode probes does require the user to employ some of the diagnostic capabilities of the MP-917 to investigate the signals delivered by the probe to be sure that the automatic measurement of the round-trip propagation time will be successful.  This application note describes this diagnostic process which will allow the user to determine if the single-diode probe will work in the target soil conditions, and how to set some of the MP-917 configuration parameters to accommodate some variation of soil conditions.  An appendix describes how to fabricate one type of single-diode probe. If several probes are to be fabricated, it is desirable that they all have similar characteristics so that one set of instrument configuration parameters will serve for all of the probes. A second appendix describes a calibration procedure which can be used to evaluate the variability to be expected among different probes. Also, some mechanical requirement for the probe (such as non-ideal location of the diode, or non-contact with soil for some portion of the probe) may require determination of the probe calibration parameters as shown in the appendix.

Connection of Probe to MP-917tc "Connection of Probe to MP-917"\l§

The probe is assumed to have an inexpensive, readily-available connector such as a type F, or no connector at all but a pigtail of 50-ohm coaxial cable.  In either case, to make the connection of the probe to the standard MP-917 cable, an adapter which combines the diode control line and the RF signal can be used as shown below, and requires a BNC connector from the adapter. The total length of the cable from the instrument front panel to the probe top should not exceed 4 meters.

µ §

Figure 1. Connection of MP-917 to Single-Diode Probes

Diode Waveformtc "Diode Waveform"\l§

To verify that the probe assembly is connected properly, examine the difference function for the reflection from the site of the diode. For a detailed description of the procedure to follow, refer to the MP-917 manual for the View·Point descriptions on obtaining diode difference functions--FINE scans.

A brief summary of the operations follows:

·	Select DIODE (F5)   - ac 0

·	GAIN (F6)   1

·	COARSE OFFSET (F4)   50 (may vary with instrument-to-probe cable length)

·	BASE  (F7)  140

·	and then FINE scan (F2)

The waveform obtained should be similar to the top waveform in Figure 2. There should be a flat baseline, then an abrupt rise with at least 7 samples (dots).  Details of the waveform following the rise will change according to the conditions the probe is in, but are of no concern for the determination of the round-trip time to the beginning of the segment, T1. The automatic controller software will determine T1 as the intersection between the extension of the baseline and a line fit to the rise--as shown in Figure 3 by the lines drawn in by hand.

µ §

Figure 4. Diode Difference Function and RAW Waveform

Open Reflection Waveformtc "Open Reflection Waveform"\l§

Refer to the MP-917 View·Point command reference section on RAW waveform scans for a detailed description of the procedure to obtain the all-diodes-off waveform which will include the open end reflection. An abbreviated description of the procedure follows.

The MP-917  raw waveform amplitude versus time is shown in the bottom trace on the diagram in Figure 5.  This display mode is enabled by selecting DIFFOFF (ALT F5).  The scan is obtained with RAW (ALT F1).  Other parameters are set similarly to the diode reflection difference function, except that the baseline offset--BASE, (F7)--will usually need to be adjusted to get the interesting parts of the waveform on-screen. Start with 180 as shown, repeatedly scan (ALT F1) and adjust the offset (F7) until the waveform is on the screen without clipping along the bottom. The value so determined for BASE will be needed for entry into the 'Probe Defaults' table as described below.

The section of this RAW waveform after T1 corresponds to the probe.  Variations to the left of T1 are of no concern (attributable to the connector, adapter reflections). Here, for successful operation with the single-diode probe, the open reflection must be easily identified.  The MP-917 will locate the minimum in the waveform, then fit a line to the rise following the minimum, and determine T2 as the intersection of the fitted line and the horizontal line through the minimum as shown in Figure 6 by the hand-drawn lines.  If a distinct rise with at least 7 points is not present, the automatic controller software will be unreliable. The waveform shown is for damp coarse sand and is therefore nearly ideal.  Variations which might be expected for other soils might include a deeper drop to the minimum, bumps along the probe portion, and a smaller amplitude rise at the open end.  These variations would require adjustment of the parameters to accommodate. If the open-end reflection transition is smaller than about 1/3 to 1/2 of the screen height, the GAIN (F6) should be adjusted upward, but avoid clipping of the waveform at the minimum or clipping of the top of the open-reflection rise (that is, operation would not be seriously affected if portions of the rise well after the open-reflection ramp are clipped on the top). To help sort out causes of various reflections in the RAW waveform display, it may be helpful to place a knife-edge short at the probe tip then obtain a RAW scan. With the short present, there will be a sharp drop at the time corresponding to the position of the short. This short waveform can be kept as the reference for comparison with the open reflection which should show a distinct rise at the same time as the short.  This comparison procedure is described in more detail in the calibration appendix.

The round-trip propagation time, T2-T1, can be approximately determined from these diagrams by observing that the horizontal time scale is given in delay counts, the lines are each 255 counts total length, and the counts-to-time conversion factor for a particular MP-917 is given in the "U0 Delay Generator Calibration Factor in psec" parameter visible in the Instrument Defaults menu. For example, for the instrument used to obtain Figure 7, the U0 Calibration factor was 141, the distance from the start to T1 is measured with a rule as 25.5 mm and the distance from the start to T2 is measured as 35.5 mm on the page.  The total length of the scan is measured as 54.5 mm. Thus:

µ §

and:

46.8 counts = 46.8×0.141 nsec/count = 6.599 nsec

These values agree with the (separate measurement) automatic-operation determination in the same soil conditions shown in the automatic-operation Figure 8.

Setting 'Instrument Default' Parameterstc "Setting 'Instrument Default' Parameters"\l§

See the MP-917, View·Point command reference description sections for more details about how to display the 'Instrument Defaults' menus.  Briefly, from the graphics display select DFLTS (F9), then select 'Instrument Defaults'.

Most of the parameters shown on the following three (screen) pages of the Instrument Defaults menu should not be altered from their factory-default values. Incorrect settings can render the MP-917 non-functional, so exercise care to avoid accidentally changing parameter values. The parameters which may require variation for use of the single-diode probes are the 'Automatic Gain Control' on menu page 1, and 'Step Ahead Counts...' on menu page 2.

The automatic gain control is usually enabled.  However, if the RAW waveform shows a small open-end reflection followed by a large rise (as the figure above does, although in this case it is not too severe) then improved performance may be obtained by disabling the AGC.  The AGC acts on the value of the waveform near the end of the scan--at the count specified in the "AGC Maximum Value" parameter--which may not represent the maximum signal of interest (the 'top' of the rise due to the open reflection). With the AGC disabled, the gain values in the Probe Defaults would need to be selected to display large enough waveforms--about half scale for the rise portion, as experimentally determined by examining the RAW waveform and adjusting the gain, as outlined in the previous section.

The 'Step-Ahead Counts...' parameter will probably need to be selected for each length of probe, and may need to be changed for a given (short) probe in highly variable soil conditions. A value is chosen to cause the automatic search to start beyond any false minima just after the diode, but small enough to avoid missing the minimum before the open reflection for the driest soil.  The selection is most delicate for very short probes in dry soil. A suggested starting value can be calculated from the theoretical, round-trip propagation time for the probe in air (shortest possible time).  For example, for a 30 cm probe and the U0 calibration 141 psec/count

µ §

A trial-and-error selection of the parameter is made by first examining the RAW signal with the probe in the soil to be measured, note the detailed shape of the waveform from the minimum to the open-end reflection. Then, be sure the automatic-operation display shows the line fit to the open-end reflection by observing the same waveshape with the fitted line drawn on the open-end rise, as shown in the typical automatic-operation waveforms Figure 9, in the right-hand waveform. The horizontal scale on this automatic-operation display is generally not the same as it is on the RAW waveform display.

INSTRUMENTS DEFAULTS MENU (PAGE 1)

1995/07/13 10:29:29

Exit Instrument Default Menu

#274b: 50

Delay 0 Value to obtain maximum reading (valid range 0 - 255):255

Delay 1 Value to obtain maximum reading (valid range 0 - 255):255

Averaging time for each measurement in msec  (valid range 0 - 255): 40

#278b: 10

#279b:  1

Time to wait after power on in Seconds (valid range 0 - 25.5):  1.0

Settling time before each scan sample in msec    (valid range 0 - 255):  8

Automatic gain control (0=Enable; not 0=Disable):  0

AGC Maximum value (valid range 0 - 255):200

#284b:255

U0 Delay generator calibration factor in psec    (range 0 - 65535):148

SquareRootOF(Dielectric constant of water) - 1 (range 6.0 - 9.0):7.962

#289i:1000

#291b:  1

#292b:  1

#293i:9600

#295b:255

INSTRUMENTS DEFAULTS MENU (PAGE 2)

1995/07/13 10:30:13

Stop fit at this number of points regardless of R^2 (valid range 0 - 255): 20

Threshold step back for fit restart (valid range 0 - 255):10

Scan threshold scaling factor (valid range 0 - 255):  5

Maximum number of times to retry fit (valid range 0 - 255):  3

#300b:  5

#301b: 10

Analog power flag. 0 means leave power on (valid range 0 - 255):255

#303b: 80

Message level (valid range 0 - 4):  2

Startup mode (valid range 0 - 7):  0

ID Number of instrument currently connected to PC:000000AEE44A

R^2 that must be exceeded to stop fit, gain=1 (0.000 < R^2 < 0.999):0.980

Smallest number of points for fit to gather, gain=1 (valid range 0 - 255):  7

Averaging time for each measurement in msec, gain=1 (valid range 0 - 255): 20

Size of wings on either side of fitted line, gain=1 (valid range 0 - 255):  5

R^2 that must be exceeded to stop fit, gain=2 (0.000 < R^2 < 0.999):0.980

Smallest number of points for fit to gather, gain=2 (valid range 0 - 255):  7

Averaging time for each measurement in msec, gain=2 (valid range 0 - 255): 20

Size of wings on either side of fitted line, gain=2 (valid range 0 - 255):  5

R^2 that must be exceeded to stop fit, gain=5 (0.000 < R^2 < 0.999):0.960

INSTRUMENTS DEFAULTS MENU (PAGE 3)

1995/07/13 10:30:23

Smallest number of points for fit to gather, gain=5 (valid range 0 - 255):  7

Averaging time for each measurement in msec, gain=5 (valid range 0 - 255):30

Size of wings on either side of fitted line, gain=5 (valid range 0 - 255):  7

R^2 that must be exceeded to stop fit, gain=10 (0.000 < R^2 < 0.999):0.960

Smallest number of points for fit to gather, gain=10 (valid range 0 - 255): 10

Averaging time for each measurement in msec, gain=10 (valid range 0-255): 40

Size of wings on either side of fitted line, gain=10 (valid range 0 - 255):  7

Baud rate select for mode 3 & 6 (0:9600, 1:4800, ..., 7:75):  0

Step-ahead counts from diode to start open search (valid range 0 - 255): 10

Datalogging Mode Interval Timing Parameter:  0

#337b:  9

Use_Select Flag (plug or memory):MEMORY

#338b:  0

#339b:  0

#340b:  0

Probe Type Currently used (valid range 0 - 10):  9

Setting 'Probe Default' Parameterstc "Setting 'Probe Default' Parameters"\l§

PROBE TYPE 9 DEFAULTS MENU

ONE SEGMENT PROBE

1995/07/13 10:30:42

Exit Probe Type 9 Default Menu

Offset for diode 0- (valid range 0 - 255):140

Gain for diode 0- (valid range 0,1,2,5,10):  1

Diode Sense (Normal(0-) or Inverted(0+)):Normal

Offset for open end (valid range 0 - 255):160

Gain for open end (valid range 0,1,2,5,10):  1

(Ts)/(Tair) (0.000 < (Ts)/(Tair) < 10.000 ): 1.000

Length of segment in mm (valid range 150.0 - 1000.0): 400.0

A factor for segment (valid range 0.000 - 10.000 ):0.000

B factor for segment (valid range 0.020 - 5.000 ):1.000

To obtain the probe defaults menu shown above, refer to the MP-917, View·Point manual command reference section. Briefly, select DFLTS (F9), then 'probe calibration defaults', then 'Type 9 probe defaults'. The value to be entered into each of the probe parameters is desbribed briefly here, and in more detail in the Appendix 'Probe Calibration'.

The Gain and offset parameters for the diode will usually be 1 and 140 respectively.

The 'Normal' sense of the diode polarity is negative (negative with respect to coaxial cable shield turns ON the diode). If the diode at the top of the probe is reversed, then the 'Diode Sense' parameter may be changed.

The Gain for the open end will usually be 1. Soil conditions may require a higher gain to increase the size of the open-end reflection, but be sure spurious minima and following rise are not thereby increased to the point where the controller regards them as open-end reflections.  Conductive soil or/and longer probes will usefully employ gains of 2 or 5. The gains set in the probe defaults table are minimums when the AGC is enabled--if the AGC system determines that the signal is too small the gain will be adjusted upward, but it will not reduce the gain below the value specified in this table.  When AGC is disabled, the gain value used during automatic operation is specified in the probe defaults table.

The Offset for the open end should be set to the value obtained experimentally, as described above, which places the minimum and open-end rise on-screen.  The value is usually not too critical and the effect of using 160(shown in the parameters above) instead of 180(used in the open-reflection display) is seen below in the automatic-operation display where the right-hand waveform is displaced down nearly to zero.

Ts/Tair = 1.55 for typical agricultural soils.  This parameter is a characteristic of the soil and is usually determined by independent means. (Refer to Moisture·Point Technical Brief, Profiling Probe Calibration)

The length of the probe over which time is measured is the length from the probe tip to the location of the diode, all of which should be exposed to the soil being measured for moisture.  If the diode location is not very near the top of the soil being probed, then the length of the soil-covered section may be specified in the probe defaults table and a corrected A value can be used--see probe calibration appendix.

Factor A = 0 represents near-ideal probe construction, where the diode is located at the soil surface.

The B probe calibration value is a correction factor which divides the measured time, it is used to adjust the measurement for incomplete interaction of the probe signal field with the soil being measured.

Determination of A, B, and segment length is the probe calibration procedure described in more detail in the Appendix.

A = 0 (or time offset correction...see calibration section below)

B = 1 (or interaction correction...see calibration section below)

Typical Automatic-Operation Waveformstc "Typical Automatic-Operation Waveforms"\l§

Figure 10. Typical Automatic-Operation Display

Appendix 1 - Probe Constructiontc "Appendix 1 - Probe Construction"\l§

The details of probe construction will be determined in large part by the need for ruggedness, environmental protection of the electrical conductors, convenience, cost, and so forth.  This description will just give simple, yet completely functional, examples of construction to illustrate the electrical connections required.

A Zegelin-type, 3-electrode, occasional-use, non-permanent installation probe (that is, not environmentally sealed, fragile) can be constructed as shown schematically in Figure A1:

µ §

Figure A1. Example 3-Prong, Single-Diode Probe

The diode is a PIN switching diode, such as the MOTOROLA MPN3404, which is a leaded diode that can be inserted into the terminal block.  The lead diagram for the MPN3404 is:

µ §

The rods may be stainless-steel welding rod or rod stock.  The length of the probe is chosen to span the region of soil over which the average moisture content is to be determined. Longer probes are better for accuracy, but will suffer greater signal loss in conductive soils.

On inserting the probe, the rods should be kept as near parallel as possible.  A guide should be used for longer probes--the guide is placed at the soil surface and the probe is inserted through the guide which holds the rods at a fixed separation.

.µ §

µ §

Encapsulate terminal block with silicon sealant to make connections water proof.

Figure A2. Example Two-prong, Single-Diode Probe.

Appendix 2 - Probe Calibrationtc "Appendix 2 - Probe Calibration"\l§

When moisture is calculated from the measured round-trip time, TM, the MP-917 first corrects the measured time for a probe's characteristics according to

µ §,

then the moisture is calculated using the formula (see Moisture·Point Technical Brief, Profiling Probe Calibration, for derivation of this formula)

µ §

where Tair is the round-trip propagation time for a signal along the probe in air, calculated from

µ § nsec.  Lseg is the segment length in mm, velocity of light in air is in mm/nsec

The factor, Ts/Tair, is a soil characteristic factor, Ts being the round-trip propagation time of a signal on the probe in completely dry soil. The denominator is determined by the dielectric constant of water, KW.

Consequently, calibration of the MP-917 for a measurement with a particular probe in a particular soil consists of determining each of the factors A, B, Ts/Tair, and Lseg, which are entered into the "Probe Defaults" table, and determining the denominator, which is entered into the "Instrument Defaults" table. The remainder of this section describes determining just the factors A, B, and Lseg, which are somewhat different for single-diode probes than they are for profiling probes.

If the single-diode probe is constructed so that the location of the diode is very near the soil surface when the probe is installed, as it would be for the construction techniques shown in Appendix 1, and the probe rods are surrounded by a sufficiently large volume of the soil being measured, then it's likely that no probe calibration is required. That is, A = 0, B = 1, and Lseg = length of rods in mm, as for an ideal probe.

If construction is such that the diode cannot be near the top of the active section of the probe (soil surface) then a constant, time-offset correction--the A value--and a modified segment length can be entered in the probe defaults. The A value is (mostly) the round-trip time for a signal to travel from the diode location to the top of the soil location. The MP-917 can be used to measure the time to this short. Since this correction time should be determined in a way similar to the usual measurement of time to the diode location, the time should be determined from a difference function between short and no-short. To obtain this difference function requires recording the two RAW waveforms--without the short, and with the short--and then subtracting them point-by-point to obtain the difference function. A step-by-step procedure outline to determine the A value is:

·	turn on logging (ALT F9)

·	record the diode waveform (as described in 'Diode Waveform')

·	record the RAW waveform for the probe in air

·	place the short at the real start of the segment, record the RAW waveform

·	turn logging off (ALT F9)

·	move the waveform data recorded (in file RAWLOG.001) to spreadsheet. Program STRIPPER.EXE may be used to extract the data from the RAWLOG file.

·	calculate the difference of the two raw waveforms in the region of the short.

·	plot the two difference waveforms (diode, open-short) and locate the intersection of each rise with each baseline.

·	calculate A as time (multiply counts by U0 calibration factor) for entry in 'Probe Defaults'

To illustrate this procedure, the following diagram gives the actual dimensions of a probe which is used to obtain the required waveforms:

µ §

Figure  A3. Dimensions of Example Probe  for determination of A coefficient.

For the example probe, the diode is located within 15 to 20 mm of the usual soil boundary--where the knife-edge short was placed to obtain the short waveform. the A correction is therefore theoretically expected to be at least 2*(15/299.704) = 0.1 nsec to 2*(20/299.704) = 0.13 nsec. This theoretical value should be thought of as a minimum--the dielectric material of the terminal block, impedance discontinuities, unmeasured lead-length, and other such neglectia will increase the time.

The MP-917 is connected to the probe according to the diagram in Figure 11. The MP-917 is also connected to the computer running View·Point, and the RAWLOG (ALT F9) feature is turned ON. The diode difference function waveform is obtained according to the procedure described in 'Diode Waveform'. Then, use the MP-917 to record the RAW waveform with the probe in air, following the procedure in 'Open Reflection Waveform'.  This pair of waveforms for the example probe is shown in Figure A4.  A knife-edge short is placed across the probe rods at the terminal block, which is where the soil surface will normally be when the probe is in the soil, and a second RAW waveform is obtained. The short waveform for the example probe, together with the same diode waveform, is shown in Figure A5. The RAWLOG is closed (ALT F9), and the View·Point program is exited (SHIFT F10).

Figure A6. RAW Waveform, Probe in Air (top), Diode difference function (bottom)

Figure  A7. RAW Waveform, short at soil surface (top), Diode difference function (bottom)

The data for the three waveforms is extracted from the rawlog file using the STRIPPER program with the command:

C:\VP> STRIPPER /w RAWLOG.001 > CALSCANS.TXT

The waveform data in the file CALSCANS.TXT is in 'quote and comma delimited' format for easy import into a spreadsheet. The region of each waveform consisting of the 20 to 30 points which spans the time of the diode difference function and the location of the short is selected. The spreadsheet is used to form the difference between the two RAW waveforms over the region of interest.  The diode difference function and the calculated difference function are plotted and lines for the baseline and rise are fitted to the data points to locate the count value corresponding to the travel time to the diode and to the short.  These fitted lines can be found with the spreadsheet's curve fitting functions, or the lines can be fitted manually as they have been in Figure .

Count	Open	Short	open-short	Diode

72	74	181	-107	18

73	73	180	-107	17

74	72	178	-106	17

75	73	178	-105	17

76	75	180	-105	16

77	75	181	-106	16

78	77	182	-105	17

79	81	185	-104	19

80	85	186	-101	27

81	90	186	-96	44

82	101	185	-84	72

83	115	180	-65	103

84	124	168	-44	125

85	128	150	-22	141

86	131	136	-5	154

87	135	123	12	164

88	137	113	24	172

89	138	105	33	178

90	141	100	41	182

91	144	94	50	185

92	144	87	57	189

Figure A8. Excerpts of Data from RAWLOG file for Waveform Examples Shown in Figures A9 and A10.

µ §

Figure A11. Spreadsheet Plots of Waveform Data, "Manual" determination of T1 (diode) and Short Time

From the intersection of the lines fitted to the rise and the baselines, read off the horizontal scale the count value representing the round-trip time to the diode location--80.0--and to the short location--81.0. The difference is 1 count = 0.148 nsec (U0 calibration factor is 0.148 nsec/count for the instrument used for this calibration).

To illustrate the size of the effect of adjusting the calibration factors from their nominal values, when the MP-917 was operated with the example probe in air, the time measured was 2.619 nsec and nominal values for the calibration coefficients (A = 0, B = 1, Lseg = 400)  gave a moisture calculation of -0.002. If we substitute the newly-determined factors A = 0.148, Lseg = 385 (measured with rule) the moisture calculation would be -0.005.  Similarly, when the probe was inserted into a water tank, the time measured was 22.866 nsec and the nominal calibration values gave a moisture calculation of 0.951 and the newly-determined factors would give 0.985. If only the length is corrected, and A = 0, the calculated moisture is 0.002 in air and 0.992 in water. Given the accuracy for the MP-917 is 0.030, just getting the length right and neglecting the non-ideal location of the diode is adequate (for this probe construction). That the calculated moisture was different from 0.000 and 1.000 in the correct-calibration cases is adequately explained by the measurement to measurement variation inherent in the MP-917 operation.